## Slide 1: Intracellular Cytokine Staining (ICS) Flow Cytometry Assay
Kent J. Weinhold, Duke CFAR Director
November 6th, 2013
CFAR Director’s Meeting
Flow Cytometry Workshop

## Slide 2
Wash
Wash
Brefeldin
Monensin
Wash
lymphocyte
erythrocyte
cytokine
6 hrs
Rest
CD107
6 h
Wash
IFN-γ PE-Cy7
TNFα Alexa700
ICS Assay Overview
Perm
IC Stain
Acquisition
Analysis
Lyse/Fix
Surface
Stain
Stimulate
Thaw
1.
2.
3.
4.
5.
6.
7.

## Slide 3: Keys to success
  Panel design (McLaughlin et. al. Cytometry 73A:400-10, 2008)
  Goal – appropriately classify events as positive or negative
    Minimize spillover/spreading error
    Bright fluorophores conjugated to dim markers
    Careful use of tandem dyes
  Intracellular stain: CD3, CD8, and CD4
    Requires mAb’s made against fixed antigens
  Kinetics of functional markers must overlap
  Use appropriate protein transport inhibitor
  CD4 responses require antigen processing
Reagent Qualification
  Titer all reagents using specific measures of performance (Murdoch et. al. Cytometry 81A:281-3, 2012)
    Signal-to-noise (SN): positive median/negative median)
    Staining index (SI): positive median/negative standard deviation
    Negative median
    Negative CV
  Measure spillover
  Bridge reagent lots – especially tandem dyes
Instrument Qualification
  Optimize PMT voltages for each type of assay (Perfetto et. a. Nat Pro 1:1522-30, 2006)
  Appropriate use of controls
    Unstimulated cells – negative control for stimulation
      Endogenous responses may be present
    FMO’s – negative control for fluorescence
  Number of events acquired
    Minimum of 120,000 viable CD3+ lymphocytes (Jaimes et. al. JIM 363:143-57, 2011)
  Analysis
    Up to 50% of assay variability (Maecker et. al. BMC Immunology 6:13, 2005 and McNeil, et. al. Cytometry 83A:728-38, 2013)
    Reproducibility dependent upon Operator expertise
    Operator training highly recommended

## Slide 4: Bridging Study Shows Degradation of Tandem Dye
Green A: IFNg PE-Cy7
Green E (PE)
Testing Date: 17Feb10
Lot #: 44563
Expiration Date: 31Mar11
IFNg PE-Cy7 SS
0.144µg/mL
Green E (PE)
Green A: IFNg PE-Cy7
Testing Date: 06May11
Lot #: 02587
Expiration Date: 31Oct12
IFNg PE-Cy7 SS
0.15µg/mL
PE-Cy7 breaking
down into PE & Cy7

## Slide 5: Compensation errors create false positive CD4 CEF response
CD3
IL-2+IFNg
CD4
Compensation error &
false positive CD4 CEF response
Compensation corrected &
No false positive CD4 CEF response
CD8
CD8
IL-2+IFNg
CD3
CD4
EOLm
SA
37.4
62.2
0.448
0.0237
30.8
61.9

## Slide 6: Excluding dim CD8+ cells significantly reduces CD8 response to CEF
EOLm
SA
CD4
CD8
SA: Site Analysis; EOLm: Centralized Manual reanalysis

## Slide 7: CD3 vs. cytokine in final plot can help visualize missing CD3 dim+
EOLm
SA
SA: Site Analysis; EOLm: Centralized Manual reanalysis

## Slide 8: Duke CFAR approaches to improve assay performance
Training
  Wet workshops
    One-on-one intensive week-long training course offered by Duke CFAR
    Collaboration between Duke CFAR Immunology & Biostatistics and Bioinformatics Cores
  FlowPET
Automated analysis

## Slide 9
0.21%
0.18%
CD4 FITC
1.9%
1.65%
CD8 PerCP-Cy5.5
IFN- + IL-2 PE
Expert Gating
Manual
Cluster Gating
Automated
Duke University Medical Center
Automated gating tools can improve signal
Notes: The automated procedure seems to be at least as sensitive as manual gating for detecting cytokine positive cells for both CD3+CD4+ and CD3+CD8+ lymphocytes.
9

## Slide 10: Acknowledgements
Duke CFAR Immunology Core – Flow Cytometry
  Kent Weinold, Duke CFAR Director
  Jennifer Enzor
  Twan Weaver
  Jianling Shi
EQAPOL
  Tom Denny
  Duke CFAR Biostatistics and Bioinformatics Core
    Cliburn Chan, Duke CFAR Core Director
    Scott White